PURPOSE

This SOP outlines general decontamination methods for refrigerators and freezers used to store Risk Group 1 (RG1) and/or Risk Group 2 (RG2) biological or biohazardous materials (worked with at ABSL2+/BSL2+ or lower).

Scope

Refrigerators and freezers used to store RG1 and/or RG2 biological or biohazardous materials must be decontaminated when they are:

- part of repairs by UF facilities or a contractor,

- part of a laboratory decommissioning process,

- being removed from service and/or stored (temporarily or permanently),

- leaving the University of Florida,

- moving to another laboratory or building on any UF Campus, or

- being released to UF Surplus Property or otherwise disposed of properly.

This SOP applies to UF laboratory personnel decontaminating refrigerators and freezers that have come into contact with up to RG2 biological or biohazardous materials. This SOP does not apply to units used to process radioactive materials or RG3 materials. Please contact EH&S Radiation Safety Team for guidance with radioactive materials. Please contact EH&S Biosafety Team for guidance with RG3 materials.

Definitions

Biological Hazards: A biological hazard, or biohazard, refers to any biological material that presents a risk or potential risk to the health of humans, animals, or the environment. Biological hazards include:

- pathogenic agents such as bacteria, viruses, fungi, and parasites,

- recombinant and synthetic nucleic acid molecules as defined by the NIH Guidelines for Research Involving Recombinant or Synthetic Nucleic Acid Molecules (NIH Guidelines),

- acute biological toxins (those with an LD50 ≤ 100 µg/kg body weight),

- human blood/tissues/organs/cells and certain human body fluids.

Responsibilities

For EH&S, PI, and laboratory personnel general responsibilities see here: Environmental Health and Safety – Regulation and Policy Hub (ufl.edu) and EH&S Biosafety Manual found here: EH&S Biosafety Manual

Area Specific SOP Author: add any additional responsibilities specific to this SOP here

Hazard Identification & Control Measures

Area Specific SOP author: Include information on how to handle a particularly hazardous substance or experimentation using a certain piece of equipment. Instructions might include recommended hazard control measures, PPE, chemical transportation, and storage. Describe transport, receiving, and storage requirements. Include secondary containment, transport devices (carts, carriers, etc.), segregation requirements, any special temperature or atmospheric requirements, and container compatibility requirements.

A. Potential Hazards

- Chemical Hazards (CH): Click or tap here to enter text.

- Biological Hazards (BH): Click or tap here to enter text.

- Physical Hazards (PH): Click or tap here to enter text.

- Electrical Hazard (EH): Click or tap here to enter text.

- Other Hazards: Click or tap here to enter text.

B. Engineering Controls / Administrative Controls

Engineering: Handwashing sinks, PPE

Administrative: SDS/MSDS Sheets available to personnel, equipment manuals are available to personnel, laboratory-specific SOPs available to personnel

C. Personal Protective Equipment (PPE)

Note: while not PPE, long pants and full coverage shoes are required to work in the laboratory.  When PPE is in place there will be no skin exposed from the neck to the floor.

  - Laboratory coat/gown

  - disposable (latex/nitrile) gloves

  - safety glasses/goggles

Materials

PPE as outlined in Hazard Identification and Control Measures

Drain pans/buckets to collect water and ice

Dustpan and scoop to sweep out any loose ice or frost

Disinfectants:

- Freshly diluted 5,000 ppm bleach solution or EPA approved tuberculocidal disinfectant

- 70% Ethanol – Ethanol evaporates too quickly to be an effective surface decontaminant.  Use this to submerge items for disinfection or following another disinfectant to prevent corrosion

Distilled Water

Non-plastic-backed paper towels

Lint-free wipes (e.g Kimtech) for delicate equipment and components

Equipment

- Computer – access to SOP

LABORATORY-SPECIFIC SAFE PRACTICES

Please outline your safe practices here.

Procedure

Note: When disposing of a refrigerator or freezer, any refrigerant (e.g. freon) in the unit must be removed before the unit can be disposed of. Before using a freshly diluted 5,000 ppm bleach solution to neutralize biological hazards, verify that there are no additional hazards present that will interact adversely with bleach.

- If applicable, disable any monitoring systems to prevent false alarms.

- Don appropriate PPE.

- Remove all items from the unit to be decontaminated and transfer them to appropriate storage.

- Following the manufacturer’s instructions, shut down and unplug the unit.

- Sweep out any loose ice and frost into the buckets and discard to a sink.

- Place towels, drain pans or buckets under shelves and doors in the unit to catch any dripping water.

- Allow the unit to defrost. Do not use sharp tools to chip off ice!

- Once the unit is completely defrosted and dry, remove any visible contamination by wiping down with an appropriate disinfectant.

  - Decontaminate collected water/ice melt as described in the Liquid Waste section of the Determination, Accumulation, Inactivation, and Disposal of Biological and Biohazardous Waste SOP.

  - Decontaminate absorbent materials (e.g.Non-plastic-backed paper towels) as described in the Solid Non-Biomedical, Biological Waste section of the Determination, Accumulation, Inactivation, and Disposal of Biological and Biohazardous Waste SOP.

- Wipe down the interior with an appropriate disinfectant (e.g. a freshly diluted 5,000 ppm bleach solution). Allow 10 min contact time.  Surfaces must remain visibly wet.

- Wipe down with distilled water or 70% ethanol to prevent corrosion.

- Log into the EH&S Admin portal at https://apps.ehs.ufl.edu/admin-portal/ and select “Equipment Decontamination Forms”.

- Create a new form by clicking “New Form” in the upper right corner of the page.

- Fill out the form as indicated. Have information regarding the decontamination and equipment ready, including the UF Asset Tag and the equipment’s manufacturer and serial number. Select all decontamination types that apply.

- Once complete, the form will be sent to the PI for verification and then to the EH&S office for approval.

- When approved, print a copy of the decontamination form and affix it to the decontaminated equipment. Remove all hazard warning stickers except radiation stickers at this time. Only EH&S Radiation Safety personnel may remove radiation stickers.

- If disposing of the unit, you may now contact a UF-approved service to remove refrigerants from the unit.

Emergency Response (Spill, Exposure, & Accident Procedures)

See lab specific spill, exposure, and accident SOPs.

References

List resources that may be useful when performing the procedure, for example, Admin policies, standards, etc.

- Meechan PJ, Potts J (2020) Biosafety in microbiological and biomedical laboratories 6th edn. U.S. Department of Health and Human Services, Public Health Service, Centers for Disease Control and Prevention, National Institutes of Health. HHS Publication No. (CDC) p 476 – 478

- University of Florida Biosafety Manual: https://www.ehs.ufl.edu/wp-content/uploads/Biosafety-Manual.pdf

- UF EH&S Equipment Decontamination https://www.ehs.ufl.edu/departments/research-safety-services/chemical-and-lab-safety/equipment-decontamination/

Documents and attachments

List applicable forms that are required to be completed in the SOP. Attach any documents used in support of the SOP, e.g., flowcharts, work instructions, pictures or diagrams, forms, and labels.

UF EH&S Equipment Decontamination Form https://www.ehs.ufl.edu/departments/research-safety-services/chemical-and-lab-safety/equipment-decontamination/